Institute of Space

Technology Islamabad

Procurement Department

RFQ-2425-00074-L

24-11-2024

To:    _____________

__________________

Subject: Invitation of Tender for Automatic Mounting Machine, Automatic Table feed Precision Cutting Machine, Grinding/Polishing Machine, Bath Sonicator, -Portable Polisher/Grinder

Dear Sir,

We invite you to quote for supply of the item as detailed in Annex-A & B enclosed. Offers must be submitted on company’s letterhead with complete specifications of offered equipment/item and undertaking given at the end of Annex-A & B along with format for bid (Annex-C & D) to be duly filled.

The terms and conditions of this tender are as under :-

- The Bids comprise a single stage, containing Two separate sealed envelopes are required to be delivered at IST-Islamabad by 12:00 Hours on 11th Dec. 2024. Tenders will be opened on same date at 12:00 Hours in the presence of bidders who desire to attend. Each envelop shall contain separately the financial proposal and the technical proposal. The envelopes shall be marked as “FINANCIAL PROPOSAL” and “TECHNICAL PROPOSAL” in bold and legible letters to avoid confusion. Initially, only the envelope marked as “TECHNICAL PROPOSAL” shall be opened. The envelope marked as “FINANCIAL PROPOSAL” shall be retained in the custody of the procuring agency without being opened. The financial proposals of bids shall be opened publicly at a time, date and venue announced and communicated to the bidders in advance.

- Sealed tenders addressed to the Deputy Director (Procurement), Institute of Space Technology, Near Rawat Toll Plaza, Islamabad Highway, Islamabad may be dropped in the tender box or mailed through courier, registered post; but it must be ensured that the same reaches our office well in time as late offers will not be considered. The envelopes should be marked clearly with the following information.

Tender No 	:     RFQ-2425-00074-L

Due Date	:      11th Dec. 2024

Tender For	:	Automatic Mounting Machine, Automatic  Table feed Precision Cutting Machine-Grinding/Cutting Machine, Bath Sonicator, -Portable Polisher/Grinder

- Rates: 	The bidder should quote their rates for :-

- Free Delivery at IST inclusive of all Govt. taxes. Details of such taxes to be provided.

- All charges within Pakistan till delivery (including port clearances, custom duties and all taxes), installation and commissioning in IST should be included and mentioned in the quotation.

- The bidder should indicate their NTN NO. & Sales Tax registration number.

- Items must be quoted as per user requirements/sample available at IST - Store.

- Before submitting any offer, bidder must consult the concerned user regarding any query/specifications or checking of sample. No change in the offer will be accepted later on.

- Offer shall be made as per our demanded specs or equivalent.

- All optional accessories are to be quoted separately as well. Essential accessories required for operational/ running of machine must be provided along with the offer.

- The organization reserves the right to modify or reduce the Scope of Work at its sole discretion, based on prevailing needs and the availability of resources or funding.

- BID SECURITY: Technical offer must be accompanied with 03% of total bid value (inclusive of accessories) in the form of pay order/ call deposit/ Demand draft/  unconditional bank guarantee from scheduled bank having 150 days validity in favor of IST-Islamabad. No offer will be entertained without bid security. The bid security will finally be adjusted/ refunded on completion of the contract. At any stage if the purchaser feels that the bidder has provided wrong information or the samples are not provided as per the specifications mentioned in their offer, the bid security will be forfeited.

- Tender fee of PKR. 2,000/- in form of DD/PO must be submitted along with the offer in favor of IST-Islamabad.

- IST FTN #: 9010613-0.

- Delivery: Item may be delivered within 120-days after issuance of Purchase Order. Successful delivery along with offloading of the ordered item at the Purchaser’s premises will be Supplier’s responsibility.

- Validity: The offer shall remain valid for our acceptance for 120-Days from the date of opening of tender. Offers with validity less than 120-days will be liable for rejection.

- Inspection: The inspection shall be carried out at IST by IST inspection committee and the inspection report (findings) shall be binding on you.

- Taxes:

- GST/CED will be paid by the PURCHASER to the SUPPLIER’S on production of a separately numbered “Tax Invoice” showing the amount of sales tax and other details under section 23 of Sales Tax Act-1990.

- As per in vogue GST SRO, 20% of the 18% GST amount will be retained by the Purchaser and later on submitted in the CBR/Government Treasury, where as 80% of the 18% GST will be paid to the Supplier for onward deposition in FBR/Government Treasury.  However, GST rate is subject to change as and when notified by Government.

- Income tax will be deducted as per rule.

- Any change in taxes levied by Govt. after opening of bids will be on account of bidder. No change in offer will be accepted later on.

- Payment:  All payments will be made through crossed cheque in favor of the SUPPLIER in Pak Rupees. The payment mode will be “100 % Payment after Delivery” and after satisfactory performance in every aspect and inspection. Advance payment is subject to provision of equal amount of Bank Guarantee valid for 60 days from date of delivery.

Note: Any representation for the delay of payment will not be accepted, until complete required documents have not been received. (Countdown starts when Store receives complete documents).

- Late Offers: The quotations received late (after opening the tender) will not be considered. Cutting/ over writing are not allowed, unless/ until authenticated under proper signature of the authorized person.

- Performance Bond: 10% of total order value to be provided by firm in shape of unconditional and irrevocable Bank Guarantee/ Demand Draft of Pakistani Scheduled Bank to be valid till the expiry of Guarantee/Warranty period within 10 working days of signing of contract. Else 0.01% of total order value for each day may be charged maximum upto 10% of contract value. Add Dir Admin & Proc decision in this regard shall be fix and final.

The release of 10% Performance bond shall occur subsequent to the termination of the validity period and upon the receipt of a written request by the firm.

- In case of any arithmetic error unit price shall prevail.

- Publication/Literature/Catalogue in Hard Copies/CDs will be provided along with offer.

- IST reserves the right to increase / decrease the quantity as per requirements.

- Only technically qualified bidders will be informed for financial opening of their bids. Firms, after qualifying technical criteria and having lowest quoted value shall be considered successful. Financial offer of technically disqualified supplier will be returned unopened along with their bid security.

- Warranty: Local agent may arrange OEM warranties of offered equipment for at least one year from the date of successful installation, commissioning.

- Bids will be evaluated in a manner that the item will bring “value for money” to the organization. Offers with lower specifications than demanded specs will be liable for rejection.

- If IST deemed necessary, the lowest bidder may be asked to provide sample before awarding of any contract/ order. In case the sample of lowest bidder is rejected, sample from other bidders will be called and contract will be awarded to technically qualified lowest bidder.

- IST may ask the lowest bidder for any clarification/providing of documents for satisfactory evaluation of offer.

- Those firms are not eligible to participate in the tender whose deliveries of stores are pending/ delayed for the five months or more or their previous performance is not satisfactory, against any IST purchase order. If participated in tender their offer may be accepted/ rejected based on previous performance.

- Conditional bids may be liable for rejection. Offers with incomplete documentation will be rejected.

- A being lowest in financial bid does not mean that the contract/ purchase order will be awarded. Following factors may also be taken into account before award of purchase order.

- Previous performance of supplier against any IST order.

- Reliability.

- Technical/ Financial capabilities.

- Acceptance of terms & conditions of tender.

- IST may ask the lowest bidder to provide a “Certificate” that firm has never been black listed which will be submitted on stamp paper of valuing Rs. 20/-

- Failure & Termination: Incase the supplier / contractor failed to deliver the stores within specified delivery period, IST shall have the following options:-

- To recover liquidated damages @ 2% of total purchase order value per month or part thereof will be imposed. The maximum limit of liquidated damages will not exceed 10% of contract value. Such damages shall be recovered from the bills, bid security and or by withholding any amount due to be paid against this or any other purchase order.

- To purchase from elsewhere, without notice to you but at your risk and cost, the stores not delivered within the stipulated period and to recover excess value so paid by us from your bills or from your security and or to cancel the contract at your risk and cost.

- Blacklisting: If the supplier failed to execute order as per laid down terms and conditions in the tender/order or the performance is not satisfactory or are found to be indulging in corrupt or fraudulent practices, the action for blacklisting of firm/ supplier/ contractor will be taken as per PPRA rule 19. In case of any dispute, the case will be referred to Vice Chancellor (IST), whose decision will be final.

- Both technical and financial offers must be submitted in two separate sealed envelopes. Each envelop should contain the documents as under.

- Supplier should be authorized distributor/reseller of OEM

| Technical Offer: - Technical quotation on letter head without prices. - Format of bid as per Annexure-C (attached). - Bid security be submitted as follows; [03 % bid security (Photocopy with values hidden] 4. Rs.2,000/- as Tender Fee. | Financial Offer: - Quotation on letter head with prices. - Format of bid as per Annexure-D (attached). |
| --- | --- |

- IST reserves the rights to accept or reject any or all tenders as a whole or in part as per PPRA Rule 33 (1). The decision in this regard will be firm final and binding on all bidders.

(Raees Ahmad)

Additional Director (Procurement)

Institute of Space Technology, Islamabad

Annexure – A

Technical Offer

RFQ-2425-00074-L

Date:     24.11.2024

Due On:     11th Dec. 2024

| Sr. No | Description | Qtynos | Yes | No | Alternative Offer |
| --- | --- | --- | --- | --- | --- |
| 1 | Automatic Mounting Press Preferably Model: Buehler Simplimet 4000 or Equivalent, Manufacturers from America, Europe, Switzerland Operation - Automatic Electrohydraulic - Mounting Mold Size 25mm, - 40 mm (to be quoted as optional) - Pressure 1000 – 4400psi (70 – 300bars) operational pressure range +,- 10% , Adjustable in 50psi (5bar) increment. Preload pressure function applies gradual pressure for delicate specimens 350psi (24bar) - Heat Time User Selectable time from 1-20 minutes - Heating Temperature 120° F – 420° F (50° C – 220° C) operation temperature range; Heating Power 1500 Watt@115VAC , 1900 Watt@230VAC - Cooling Time User Selectable or Smart Cool - User Selectable time from 1-30 minutes, Adjustable in 10 second increments, Smart Cool Intelligent cooling system automatically cools until sample temperature reaches 40°C - Thermoplastic Function Automatically sets correct parameters from thermoplastic resins such as Trans Optic TM and cools the mount in a controlled, linear fashion appropriate for the mounting media - Working Environment: Temperature: 40° F to 104° F , 5° C to 40° C; Humidity: 30 – 90% humidity, non-condensing - Voltage, Frequency 85-264VAC, 50,60Hz 1 phase - Operation, Service, Maintenance and Calibration Manuals: Soft and hard copy of the operation, service & maintenance manuals including activation keys, passwords; parts catalogue; error code & parameters list should be provided. Installation & training: Free of cost on-site Installation, commissioning, maintenance and troubleshooting training for at least 04 persons from faculty will be given in IST from OEM expert. Warranty: Preferably Warranty for 10 years - Availability of spares. The vendor has to guarantee the availability of attachments, spares support, repair and free of cost update throughout life of equipment. Other Conditions: Efficient support by OEM through remote access to instrument parameters and software if required. Quotation with detailed breakdown and details of optional items should be provided separately. - All charges within Pakistan till delivery in IST including port clearances and taxes should be mentioned in quote. - Organizations where the equipment or its closest model has been supplied in Pakistan; name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are required to be met as close as possible for the selection of the equipment | 1 |  |  |  |
| 2 | Grinding and Polishing Machine Preferably Model: FORCIPOL 102-250 or Equivalent, Manufactured in: USA, Europe, Japan Benchtop Grinding polishing machine body made of ABS, which is resistant to corrosion, has high impact resistance and has high electrical insulation properties. The technical values of the machine shall be as follows: - Fully Automatic Polishing Device with single disc, the rotation speed is infinitely variable between 50 and 600 rpm. The motor power should be1 HP or more., 230V 50 Hz., 1 Phase. Disc diameter will be 250 mm Ø. Motor over speed protection relay. - The disc diameter of the Sanding Polishing Device is 250mm. - There should be a control unit on the device that enables the use and control of the disk speed, processing time, cooling water and start/stop functions of the devices. The time parameter must be programmable. There should be an emergency stop button on the front panel of the device. - The device should have a water saving feature. The device must have a water activation system that can be selected automatically or manually. When the water is left on, the water should be turned off within 2 minutes. - Should have water spray nozzle to prevent the accumulation of sawdust in the water tank. - No splash of water and consumed material to the environment. Polishing device reservoir should be easy to clean with its removable water system. - The device should have the feature of washing and drying the disc after one-touch operation. - Machine should be designed to work without vibration. Following accessories should include in standard package and to be supplied with the equipment. - One 250 mmØ PVC disc - One 250 mmØ special water retaining ring - One 250 mmØ sanding ring - One cover - Complete with water regulating valve, water inlet and outlet connection hoses. Following consumables should include as OPTIONAL and should be quoted with detailed breakdown. - Packs of 100 pieces EACH of 250 mmØ sandpaper in 180 grit, 220 grit, 320 grit, 800 grit, 1200 grit,1500 grit, 2000 grit - 250 mm Ø polishing pads, Qty -100 pieces (in packs of 10) of the EACH of: 6 microns 39-015-250 FEDO 6J, 1microns (10 Packs) 39-067-250 FEDO 1M, - 10 piece of 250 ml. Diamond suspension of: 6 micron 39-430-M, DIAPAT-M and 1 micron 39-410-M, DIAPAT-M - Operation, Service, and Maintenance Manuals - Soft and hard copy of the operation, service & maintenance manuals including parts catalogue; error code & parameters list should be provided. Installation & training Free of cost on-site Installation, commissioning, maintenance and troubleshooting training for at least 04 persons from faculty will be given in IST from OEM expert. Warranty: Preferably for 10 years Availability of spares. The vendor has to guarantee the availability of attachments, spares support, repair and free of cost update throughout life of equipment. Other Conditions - Efficient support by OEM through remote access to instrument parameters and software if required. - Quotation with detailed breakdown and details of optional items should be provided separately. - All charges within Pakistan till delivery in IST including port clearances and taxes should be mentioned in quote. - Organizations where the equipment or its closest model has been supplied in Pakistan; name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are desirable for the selection of the equipment | 1 |  |  |  |
| 3 | Bath Sonicator Preferably Model: (KEMET 3P 1.9L or Equivalent) Made by: KEMET UK, 1)ULTRASONIC FREQUENCY: 37KHZ 2) INTERIOR MATERIAL: STAINLESS STEEL 3) CAPACITY: 1.9L 4) TIME CONTROLLER: TIMING CONTROL RANGE 0 TO 30 MINUTES 5) POWER SUPPLY: AC 100~120V 50,60HZ AC 220~240V 50,60HZ 6) ULTRASONIC POWER: HEATING POWER: nom,peak (W): 80,320 7) Heating power (W): 200 Operation, Service, Maintenance and Calibration Manuals. Soft and hard copy of the operation, service & maintenance manuals Installation & training Free of cost on-site Installation, commissioning, maintenance and troubleshooting training for at least 04 persons from faculty will be given in IST from OEM expert. Warranty preferably for 10 years Availability of spares. The vendor has to guarantee the availability of attachments, spares support, repair and free of cost update throughout life of equipment. Other Conditions • Efficient support by OEM through remote access to instrument parameters and software if required. • Quotation with detailed breakdown and details of optional items should be provided separately. • All charges within Pakistan till delivery in IST including port clearances and taxes should be mentioned in quote. • Organizations where the equipment or its closest model has been supplied in Pakistan; name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are desirable for the selection of the equipment | 1 |  |  |  |
| 4 | AUTOMATIC TABLE FEED PRECISION CUTTING Machine: Preferably Model : MICRACUT- 202 AX or Equivalent Made by: Turkey,EU Device; It will be used for cutting /sectioning a wide variety of materials such as metal, ceramic, composite, bio-materials, dental material, electronic components, crystals, minerals, glass. - To operate at 220V 50 Hz single-phase voltage. - The device must be PLC controlled. - The device should have a programmable color and touch 7” inch LCD screen. - Max. A cutter disc with a diameter of Ø 200 mm can be mounted. Cutting disc speed 300-4000 rpm. will be infinitely adjustable. - Strokes along Z, Y, and X Axes, should be 45 210 and 60 mm or more; respectively. - If automatically cutting direction is along Y axis, the sample can be positioned thru motorized system along X-Z axes. - Should be able to make slices of given thickness. - Cutting feed rate 5 – 3000 micron/s ; should be adjustable. - Device should stop automatically when it reaches given cutting depth. - Table dimensions should be 365 mm X 205 mm. (when x axis is opened) - The distance between the device table channels should be 12 mm. - The device should have serial intermittent cutting method and cutting methods with intelligent feed rate control. - Cutting motor power should be 1.1 kW. - It should allow positioning with 3 axis proportional speed controlled joystick. - The device should be able to record 99 different programs. - Device lighting should be provided with LED lighting system. - The device must have a laser alignment unit. All operational, maintenance and service manuals (soft and hard copies), rights such as dongle software passwords etc shall be provided to user. - The supplier should provide details of all operational, essential spare/consumables to be required for 10 years’ smooth operations/maintenance of the equipment. Following accessories to quoted, with detailed breakdown, separately as OPTIONS - 1 Universal Sample Holder Vise - 1 vise for longitudinal section - 1 vise for round specimens - 1 pcs Cam mechanism quick vise left group - 1 clamping stand for sample holders - 1 set of Ø 75 flanges Following consumables to quoted, with detailed breakdown, separately as OPTIONS - 200mm Ø following Diamond Cutting Discs: Metal Bonded (quantity x2), for stainless steels and non-ferrous materials (quantity x20), for medium hard and hardened steels (quantity x20), Resin Bonded (quantity x20), Abrasive Cutting Discs (quantity x20), - 2 bottles of Cutting Fluid, 1 litre each Training, free of cost: - Inspection and training of two weeks - Basic training covering routine operations/ maintenance - All expenses for the trainings will be at Sellers’s charge Warranty: Supplier shall be responsible for installation, commissioning and after sale service during the warranty period and thereafter. - The warranty period shall be for at least 01 years from the date of satisfactory commissioning - Any spares needed during this period shall be provided by supplier free of charges. - The defects within the warranty period shall be rectified free of cost by arranging free replacement. Details of other optional accessories; attachments and necessary spares. Service backup support for 20 years. Name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are desirable for the selection of the equipment. | 1 |  |  |  |
| 5 | PORTABLE GRINDER/POLISHER, Manufacturers: USA, Japan, Germany, UK SPECS: To be operated with at 110V supply as well as Battery Operated (Chargeable) unit . Handle with 90° degree angle head, disc dia: 30-mm, speed: 1000 – 25000, rpm (variable in steps), Power 76w, Max. Torque: 4.5-5 Cn.M with straight cord: 1.5m, Power Supply: AC 110 V/60 Hz, digital speed indicator, clockwise / anticlockwise rotation, engine speed recognition, overload indicator, reduction gearing 01: speed: reduced to ¼ torque: 4 times increased, angle part: 90°: collet ø1.0-3.0mm, Collet Ø1.0-3.0mm Can also be operated with chargeable battery Two Separate Batteries Must Be Included In Package Preferably Warranty For 10 Years. |  |  |  |  |

Undertaking:

I/We, agree with the terms/conditions of tender inquiry. We have carefully read the terms and conditions mentioned in tender documents. We accept all these conditions.

Signature: _________________

Name: ____________________

Designation: _______________

Stamp: ___________________

(To be Printed On Company’s Letterhead)

Annexure – B

Financial Offer

RFQ-2425-00074-L

Date:     24.11.2024

Due On:  11th Dec. 2024

| Sr. No | Description | Qtynos | Yes | No | Alternative Offer |
| --- | --- | --- | --- | --- | --- |
| 1 | Automatic Mounting Press Preferably Model: Buehler Simplimet 4000 or Equivalent, Manufacturers from America, Europe, Switzerland Operation - Automatic Electrohydraulic - Mounting Mold Size 25mm, - 40 mm (to be quoted as optional) - Pressure 1000 – 4400psi (70 – 300bars) operational pressure range +,- 10% , Adjustable in 50psi (5bar) increment. Preload pressure function applies gradual pressure for delicate specimens 350psi (24bar) - Heat Time User Selectable time from 1-20 minutes - Heating Temperature 120° F – 420° F (50° C – 220° C) operation temperature range; Heating Power 1500 Watt@115VAC , 1900 Watt@230VAC - Cooling Time User Selectable or Smart Cool - User Selectable time from 1-30 minutes, Adjustable in 10 second increments, Smart Cool Intelligent cooling system automatically cools until sample temperature reaches 40°C - Thermoplastic Function Automatically sets correct parameters from thermoplastic resins such as Trans Optic TM and cools the mount in a controlled, linear fashion appropriate for the mounting media - Working Environment: Temperature: 40° F to 104° F , 5° C to 40° C; Humidity: 30 – 90% humidity, non-condensing - Voltage, Frequency 85-264VAC, 50,60Hz 1 phase - Operation, Service, Maintenance and Calibration Manuals: Soft and hard copy of the operation, service & maintenance manuals including activation keys, passwords; parts catalogue; error code & parameters list should be provided. Installation & training: Free of cost on-site Installation, commissioning, maintenance and troubleshooting training for at least 04 persons from faculty will be given in IST from OEM expert. Warranty: Preferably Warranty for 10 years - Availability of spares. The vendor has to guarantee the availability of attachments, spares support, repair and free of cost update throughout life of equipment. Other Conditions: Efficient support by OEM through remote access to instrument parameters and software if required. Quotation with detailed breakdown and details of optional items should be provided separately. - All charges within Pakistan till delivery in IST including port clearances and taxes should be mentioned in quote. - Organizations where the equipment or its closest model has been supplied in Pakistan; name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are required to be met as close as possible for the selection of the equipment | 1 |  |  |  |
| 2 | Grinding and Polishing Machine Preferably Model: FORCIPOL 102-250 or Equivalent, Manufactured in: USA, Europe, Japan Benchtop Grinding polishing machine body made of ABS, which is resistant to corrosion, has high impact resistance and has high electrical insulation properties. The technical values of the machine shall be as follows: - Fully Automatic Polishing Device with single disc, the rotation speed is infinitely variable between 50 and 600 rpm. The motor power should be1 HP or more., 230V 50 Hz., 1 Phase. Disc diameter will be 250 mm Ø. Motor over speed protection relay. - The disc diameter of the Sanding Polishing Device is 250mm. - There should be a control unit on the device that enables the use and control of the disk speed, processing time, cooling water and start/stop functions of the devices. The time parameter must be programmable. There should be an emergency stop button on the front panel of the device. - The device should have a water saving feature. The device must have a water activation system that can be selected automatically or manually. When the water is left on, the water should be turned off within 2 minutes. - Should have water spray nozzle to prevent the accumulation of sawdust in the water tank. - No splash of water and consumed material to the environment. Polishing device reservoir should be easy to clean with its removable water system. - The device should have the feature of washing and drying the disc after one-touch operation. - Machine should be designed to work without vibration. Following accessories should include in standard package and to be supplied with the equipment. - One 250 mmØ PVC disc - One 250 mmØ special water retaining ring - One 250 mmØ sanding ring - One cover - Complete with water regulating valve, water inlet and outlet connection hoses. Following consumables should include as OPTIONAL and should be quoted with detailed breakdown. - Packs of 100 pieces EACH of 250 mmØ sandpaper in 180 grit, 220 grit, 320 grit, 800 grit, 1200 grit,1500 grit, 2000 grit - 250 mm Ø polishing pads, Qty -100 pieces (in packs of 10) of the EACH of: 6 microns 39-015-250 FEDO 6J, 1microns (10 Packs) 39-067-250 FEDO 1M, - 10 piece of 250 ml. Diamond suspension of: 6 micron 39-430-M, DIAPAT-M and 1 micron 39-410-M, DIAPAT-M - Operation, Service, and Maintenance Manuals - Soft and hard copy of the operation, service & maintenance manuals including parts catalogue; error code & parameters list should be provided. Installation & training Free of cost on-site Installation, commissioning, maintenance and troubleshooting training for at least 04 persons from faculty will be given in IST from OEM expert. Warranty: Preferably for 10 years Availability of spares. The vendor has to guarantee the availability of attachments, spares support, repair and free of cost update throughout life of equipment. Other Conditions - Efficient support by OEM through remote access to instrument parameters and software if required. - Quotation with detailed breakdown and details of optional items should be provided separately. - All charges within Pakistan till delivery in IST including port clearances and taxes should be mentioned in quote. - Organizations where the equipment or its closest model has been supplied in Pakistan; name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are desirable for the selection of the equipment | 1 |  |  |  |
| 3 | Bath Sonicator Preferably Model: (KEMET 3P 1.9L or Equivalent) Made by: KEMET UK, 1)ULTRASONIC FREQUENCY: 37KHZ 2) INTERIOR MATERIAL: STAINLESS STEEL 3) CAPACITY: 1.9L 4) TIME CONTROLLER: TIMING CONTROL RANGE 0 TO 30 MINUTES 5) POWER SUPPLY: AC 100~120V 50,60HZ AC 220~240V 50,60HZ 6) ULTRASONIC POWER: HEATING POWER: nom,peak (W): 80,320 7) Heating power (W): 200 Operation, Service, Maintenance and Calibration Manuals. Soft and hard copy of the operation, service & maintenance manuals Installation & training Free of cost on-site Installation, commissioning, maintenance and troubleshooting training for at least 04 persons from faculty will be given in IST from OEM expert. Warranty preferably for 10 years Availability of spares. The vendor has to guarantee the availability of attachments, spares support, repair and free of cost update throughout life of equipment. Other Conditions • Efficient support by OEM through remote access to instrument parameters and software if required. • Quotation with detailed breakdown and details of optional items should be provided separately. • All charges within Pakistan till delivery in IST including port clearances and taxes should be mentioned in quote. • Organizations where the equipment or its closest model has been supplied in Pakistan; name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are desirable for the selection of the equipment | 1 |  |  |  |
| 4 | AUTOMATIC TABLE FEED PRECISION CUTTING Machine: Preferably Model : MICRACUT- 202 AX or Equivalent Made by: Turkey,EU Device; It will be used for cutting /sectioning a wide variety of materials such as metal, ceramic, composite, bio-materials, dental material, electronic components, crystals, minerals, glass. - To operate at 220V 50 Hz single-phase voltage. - The device must be PLC controlled. - The device should have a programmable color and touch 7” inch LCD screen. - Max. A cutter disc with a diameter of Ø 200 mm can be mounted. Cutting disc speed 300-4000 rpm. will be infinitely adjustable. - Strokes along Z, Y, and X Axes, should be 45 210 and 60 mm or more; respectively. - If automatically cutting direction is along Y axis, the sample can be positioned thru motorized system along X-Z axes. - Should be able to make slices of given thickness. - Cutting feed rate 5 – 3000 micron/s ; should be adjustable. - Device should stop automatically when it reaches given cutting depth. - Table dimensions should be 365 mm X 205 mm. (when x axis is opened) - The distance between the device table channels should be 12 mm. - The device should have serial intermittent cutting method and cutting methods with intelligent feed rate control. - Cutting motor power should be 1.1 kW. - It should allow positioning with 3 axis proportional speed controlled joystick. - The device should be able to record 99 different programs. - Device lighting should be provided with LED lighting system. - The device must have a laser alignment unit. All operational, maintenance and service manuals (soft and hard copies), rights such as dongle software passwords etc shall be provided to user. - The supplier should provide details of all operational, essential spare/consumables to be required for 10 years’ smooth operations/maintenance of the equipment. Following accessories to quoted, with detailed breakdown, separately as OPTIONS - 1 Universal Sample Holder Vise - 1 vise for longitudinal section - 1 vise for round specimens - 1 pcs Cam mechanism quick vise left group - 1 clamping stand for sample holders - 1 set of Ø 75 flanges Following consumables to quoted, with detailed breakdown, separately as OPTIONS - 200mm Ø following Diamond Cutting Discs: Metal Bonded (quantity x2), for stainless steels and non-ferrous materials (quantity x20), for medium hard and hardened steels (quantity x20), Resin Bonded (quantity x20), Abrasive Cutting Discs (quantity x20), - 2 bottles of Cutting Fluid, 1 litre each Training, free of cost: - Inspection and training of two weeks - Basic training covering routine operations/ maintenance - All expenses for the trainings will be at Sellers’s charge Warranty: Supplier shall be responsible for installation, commissioning and after sale service during the warranty period and thereafter. - The warranty period shall be for at least 01 years from the date of satisfactory commissioning - Any spares needed during this period shall be provided by supplier free of charges. - The defects within the warranty period shall be rectified free of cost by arranging free replacement. Details of other optional accessories; attachments and necessary spares. Service backup support for 20 years. Name and contact number of the end user. Services provided to the end user since it has been sold; include maintenance; calibration; supply of spares and, or attachments Note: Above mentioned specifications are desirable for the selection of the equipment. | 1 |  |  |  |
| 5 | PORTABLE GRINDER/POLISHER, Manufacturers: USA, Japan, Germany, UK SPECS: To be operated with at 110V supply as well as Battery Operated (Chargeable) unit . Handle with 90° degree angle head, disc dia: 30-mm, speed: 1000 – 25000, rpm (variable in steps), Power 76w, Max. Torque: 4.5-5 Cn.M with straight cord: 1.5m, Power Supply: AC 110 V/60 Hz, digital speed indicator, clockwise / anticlockwise rotation, engine speed recognition, overload indicator, reduction gearing 01: speed: reduced to ¼ torque: 4 times increased, angle part: 90°: collet ø1.0-3.0mm, Collet Ø1.0-3.0mm Can also be operated with chargeable battery Two Separate Batteries Must Be Included In Package Preferably Warranty For 10 Years. |  |  |  |  |

Undertaking:

I/We, agree with the terms/conditions of tender inquiry. We have carefully read the terms and conditions mentioned in tender documents. We accept all these conditions.

Signature: _________________

Name: ____________________

Designation: _______________

Stamp: ___________________

(To be Printed On Company’s Letterhead)

(To be Printed and Typed On Company’s Letterhead)

Annexure – C

RFQ-2425-00074-L

Date:      24.11.2024

Due On:   11th Dec. 2024

FORMAT FOR BID

1.	Tender Fee PO/DD No.								:		_______________________________

2. 	Tender Fee PO/DD dated							      :		_______________________________

3.	Bid Security PO/DD No.								:		_______________________________

4. 	Bid Security PO/DD dated							:		_______________________________

5.	Bid Security PO/DD issued by						:		_______________________________

(Bank Name)

6. 	Complete Technical specifications 				      :		_________________________YES/NO

& catalog attached

7.	Delivery Schedule									      :		_______________________________

8. 	Validity of Acceptance							      :		_______________________________

9.	Country of Origin.									      :		_______________________________

10. Profile of fim for performing such 				      :		_________________________YES/NO

Project. (Attached)

11.	Free of cost onsite installation/ 					      :		____________________MANDATORY

Commissioning & Training

12.	Performance Bond to be provided.				      :		____________________MANDATORY

13.	Guarantee/Warranty.								     :		___________________ _             3-Year

14.	Warranty response time and replacement time.	     :		_________________________YES/NO

15.	Owner name										     :		________________________________

Owner E-mail										     :		________________________________

Owner Contact Number							     :		________________________________

17.	Company full address								     :		________________________________

Contact Person									     :		________________________________

Cell #												     :		________________________________

E-mail												     :		________________________________

15.	Acceptance of terms & conditions of tender 		    :		      _________________________YES/NO

__________________

Signature & Seal

(To be Printed and typed On Company’s Letterhead)

Annexure – D

RFQ-2425-00074-L

Date:      24.11.2024

Due On:   11th Dec. 2024

FORMAT FOR BID

1.	Tender Fee PO/DD No.								:		_______________________________

2. 	Tender Fee PO/DD dated							      :		_______________________________

3.	Bid Security PO/DD No.								:		_______________________________

4. 	Bid Security PO/DD dated							:		_______________________________

5.	Bid Security PO/DD issued by						:		_______________________________

(Bank Name)

6. 	Complete Technical specifications 				      :		_________________________YES/NO

& catalog attached

7.	Delivery Schedule									      :		_______________________________

8. 	Validity of Acceptance							      :		_______________________________

9.	Country of Origin.									      :		_______________________________

10. Profile of fim for performing such 				      :		_________________________YES/NO

Project. (Attached)

11.	Free of cost onsite installation/ 					      :		____________________MANDATORY

Commissioning & Training

12.	Performance Bond to be provided.				      :		____________________MANDATORY

13.	Guarantee/Warranty.								     :		___________________ _             3-Year

14.	Warranty response time and replacement time.	     :		_________________________YES/NO

15.	Owner name										     :		________________________________

Owner E-mail										     :		________________________________

Owner Contact Number							     :		________________________________

17.	Company full address								     :		________________________________

Contact Person									     :		________________________________

Cell #												     :		________________________________

E-mail												     :		________________________________

15.	Acceptance of terms & conditions of tender 		     :		_________________________YES/NO

_____________

Sign & Seal